Microdetection of Chromium Mixed  – Valences using          CuNi/Ti Electrode and Linear Sweep Voltammetry

Nirwan Syarif*

Departemen of Chemistry, Universitas Sriwijaya, Indralaya, Indonesia 30814

Abstract: An electrode of non-precious metal has the potency to be utilized as a working electrode for voltammetry measurements of Cr3+ and Cr6+.  The analytical performance of the electrode in  determination specific metal species qualitatively and quantitatively was studied. The detection data were recorded and analysed using cyclic voltammetry (CV) and linear sweep voltammetry (LSV), whereas scanning electron microscope (SEM) and x-ray diffraction (XRD) spectroscopy were used to determine the surface morphology and the presence of crystalline  in the electrode. Electroanaytical performance was determined by linear sweep voltammetry. The results show that the reduction of Cr6+ to Cr3+ appeared at +0.84 V (vs Ag/AgCl) and the reduction of Cr3+ to Cr2+ at +0.74 V (vs Ag/Ag).  The optimum conditions for measuring Cr3+ were found at pH 6, deposition time of 30 seconds, and a scan rate of 160 mVs-1. The optimum pH for measuring Cr6+ is 4, deposition time of 60 seconds, and a scan rate of 140 mVs-1. The regression curve for the Cr3+ is linear in concentration 1 – 10 gL-1 with a correlation coefficient is 0.9883 and detection limit of 2.08 gL-1, while for the Cr6+ is linear in the range of 1 – 10 gL-1 with a correlation coefficient of 0.99 and a detection limit of 2.18 gL-1. There is a slight difference in the individual measurement current and the mixture of Cr3+ - Cr6+  but a good agreement for the oxidation-reduction potential. The measurement data analysis shows the feasibility of the electrode and the measurement system developed.

Keywords: alloy;  potentiostat; diffractogram; voltammogram; electrolyte.

1. Introduction

One element that is harmful to health is chromium. Chromium is a heavy metal whose presence in waters. It needs to be analyzed because chromium belongs to hazardous toxic materials. Chromium in water exists in two stable oxidation states, namely Cr3+ and Cr6+. Chromium (III) is an important species in living organisms. Chromium (III) plays a role in maintaining the normal glucose, fat, and protein metabolism in mammals 1. Chromium (VI) inversely is a  toxic and carcinogenic substance to humans, although in relatively low concentrations 2.

Several methods are commonly used for the determination of Cr, such as atomic absorption spectrophotometry (AAS) 3, coupled plasma inductive - atomic emission spectrophotometry (ICP-AES) 4, coupled plasma inductive-mass spectrometry (ICP-MS) 5. These methods require expensive costs, and they cannot distinguish the Cr species instead of directly measured Cr as total mixed species. Therefore, the mixed-valence needs to be separated into its valence for the measurement. Several voltammetric methods have been widely used for heavy metal analysis 6. Voltammetry is an electroanalytical method based on the oxidation-reduction process on the electrode surface. This method allows mixtures of ions to be measured individually in one step 7. Stripping 8  and cyclic voltammetry 9 were chosen because it has high sensitivity 10, low detection limit on the ppb scale 11, easy use and easy sample preparation 12.

The electrode is the place for electrochemical processes, i.e., a faradaic process that the electric current generates and a non-faradaic process that ionic absorbing takes place at the electrode. The working electrode is where the observed electrochemical reactions occurred. The working electrodes commonly used are metal solids 13 such as gold (Au), platinum (Pt), silver (Ag), semiconductor electrodes  14 and carbon electrodes such as glassy carbon, graphite, or carbon paste 15.

One of the electrodes used is platinum (Pt) electrochemical application. Pt electrodes are known to have the ability to reduce Cr6+ to Cr3+.  Platinum employs several forms for the analytical application of ionic Cr electrolytes16. Electrodeposition for various applications frequently used Pt wire 17, and some other such as Pt powder 18, Pt foil 19, and Pt rod for various applications 9. Pt metal plated on a support metal such as titanium is also worthy of attention because it has an excellent current density 20. Partial substitution of Pt metal is a frequent step since it is relatively costly. Wijenberg et al. reported the ability of Pt/Ir electrodes to carry out oxidation-reduction reactions 21. The electrode of IrO2 coated Ti oxidizes simulated wastewater electrochemically 22. Zn coating on Fe produced by electrodeposition technique 23. The voltammetry test results from this study indicate the presence of electrochemical activity in the electrolyte solution. Sn coated on the Fe surface is reported to be more likely to have reducing activity. These three reports provide an overview of the feasibility of using electrocatalyst metal alloys with metal supports.

Titanium is a metal widely applied as a support in preparing electrodes24. This paper reports the results of research on the manufacture of electrodes, namely Fe and Ni metal alloys on the surface of Ti. The measurement conditions, such as pH, deposition time, and scan rate, affect the electrode's linearity and the limit of detection (LOD) 25. These factors affect the ionic deposition on the electrode's surface and the performance of the measurement.

2. Experimental

## Materials

The electrodes, Ag/AgCl and Pt were bought from our local company.  The chemicals, such as deionized water,  Isopropanol (Fluka),  titanium plate were analytical grade and were used without purification. NiCl2, 6H2O, Cu2O were used as precursor for preparing electrode. CrCl3.6H2O (Merck), K2Cr2O7 (Merck), NaNO3 (Merck), EDTA (Merck), and CH3COONa  were used to prepare stock solutions. Sandapaper (Toho) was purchased from local store.

## Working Electrode

## The CuNi/Ti electrode was prepared using NiCl2. 6H2O and Cu2O  as a precursor. Both dissolved in pure isopropanol (Fluka). .An electric furnace heated the titanium plate at 400°C to have its rough surface 22. The titanium plate was sandblasted to ensure good adhesion of the deposit on its surface. The plate was cleaned using water, rinsed in detergent, and washed in water and isopropanol after heating. This step cleaned the Ti surface from residual sands from sandpaper . The surface dried in an oven at 80 C. The coating was prepared on a rough metal surface The precursor was coated to the surface using a doctor blade knife on a cleaned and coarse 20 mm x 10 mm titanium (Ti) plate. The coated surface was put in an oven for 15 min at 80°C to allow the solvent is evaporated. The layer was fired at 400°C in the furnace for 15 min to allow the decomposition of the precursor. The procedures were repeated five times to achieve the desired weight of the coating. A 2.5 g of total mass of the metal alloy used was calculated from its precursor with Cu:Ni (3:1) ratio..

## Measurement setup

Two electrolyte solutions were prepared, i.e., electrolyte content Cr3+ and Cr6+. The mixed-valence electrolyte was made from the mixture of Cr3+ and Cr6+ dissolved in one solution. A 0.514 grams of CrCl3.6H2O was weighed and dissolved in 100 mL deionized water. Standard solution of Cr3+ ion was prepared from 1000 ppm solution, which was then diluted gradually into several concentrations, namely 0.1; 0.5; 1; 5 and 10 g/L as much as 100 mL each. Cr6+ was made from K2Cr2O7 (Merck). K2Cr2O7 powder was weighted as much as 0.565 grams, then dissolved with distilled water in a 100 mL volumetric flask to the limit mark and stirred to have 1000 ppm of Cr6+ solution. Standard solution of metal ion Cr6+ was prepared from 1000 ppm solution, which was then diluted gradually into several concentrations, i.e., 0.1; 0.5; 1; 5 and 10 gL-1 as much as 100 mL each. The electrolyte solution was prepared by mixing 21.3 g NaNO3, 1.46 g EDTA, and 1.64 g NaAc, then dissolved with deionized water in a 100 mL volumetric flask to the mark.

The instrumentation configuration for determining the optimum conditions for measuring Cr3+ and Cr6+, as well as for the speciation of Cr3+ and Cr6+, is as follows: the PGSTAT204 Autolab Metrohm potentiostat is connected to the electrodes.. The other end of  potentiostat werer connected to the computer to obtain data and power source. The solution was placed in the potentiostat for cyclic voltammetry and linear sweep voltammetry tests with 3 electrodes in half cell system. The CuNi/Ti was used as the working electrode, Ag/AgCl as the reference electrode, and Pt as the counter electrode in an electrolyte solution containing Cr3+ and Cr6+.

## Determination of optimum pH

A 10 ml of 1 gL-1 Cr3+ solution with 2.5 mL of the electrolyte solution was put into a voltammetry cell with a pH variation of 4,5,6,7 and 8 then measured the current at a potential of -700 mV to 990 mV, deposition time was 30 seconds and scanned 160 mVs-1. Measurement of Cr6+ was carried out in at a potential window of -700 mV to 990 mV, deposition time of 60 seconds, and a scan rate of 140 mVs-1.

## Determination of deposition time

A 10 ml of Cr3+ 1 gL-1 solution with 2.5 ml of the electrolyte solution was put into a voltammetry cell at the best pH conditions and then measured with a current at a potential window of -700 mV to 990 mV with variations in deposition time of 0, 15, 30, 60 and 120 seconds, and at scan rate 160 mVs-1. The measurement of Cr6+ was carried out at a potential window of -700 mV to 990 mV and scan rate of 140 mVs-1.

## Determination of the Optimum Scan Rate

A 10 ml of 1 gL-1 Cr3+ solution with 2.5 ml of the electrolyte solution was put into a voltammetry cell at the best pH and deposition time and then measured the current at a potential window of -700 mV to 990 mV, with the variations in the scan rate 100, 110, 120, 140, 160, 180, 200 mVs-1. Measurement of Cr6+ was carried out in the same way as the above procedure.

## Linear Regression Plot

Determination of the regression plot were done by varying the concentration of Cr3+ or Cr6+ 0.1; 0.5; 1; 5; 10 gL-1. Measurements were carried out at optimum measurement conditions. The concentration of Cr6+ solution was prepared using the same procedure.

## Determination of the Limit of Detection

The procedure was done by measuring the concentration of Cr3+ and Cr6+, which gives a linear regression of current. The current response measurement was repeated three times. The limit of detection was obtained from the following calculation using

(1)

(2)

Where Sy/x is the standard deviation of measurement, with y is current, n is number of data, a is the slope of the linear regression curve, and LOD is the detection limit.

3. Results and Discussion

X-ray diffraction measurements were carried out to determine the metal elements composed  in the electrode. The diffractogram of the electrode 26 in Figure 1(a) shows the presence of TiNi (2Ɵ  = 26.6,  40.0, 48.8,  54.5) as described in paper27,  and overlaps with the diffraction peak of TiCu (2Ɵ = 18.15, 40.0, 54.5)  refer to the report by reference . These elements form alloys and composites of the electrode. These results indicate that the electrode is made of metal and not its oxide. SEM image (Figure 1(b) confirms the diffractogram. Figure 1 shows the SEM representation of the thermally prepared FeNi electrode morphology. The TiCu surface shows overlapping layers of TiNi, cracks, and flakes. The thermal shock that arose during the electrode removal from the oven caused the various cracks observed on the surface of the electrodes. The cracks appeared during their removal from the oven. Ti which Cu well covered, only appeared as low-intensity peaks at 2Ɵ = 40 and 45, which confirmed the presence of TiCu.

(a)

(b)

Fig. 1. (a) SEM image (b) Diffractogram of the electrode for TiCu and TiNi

The surface consists of the main structure of the electrode and crack of various sizes. The cracks are impurities that are then cleaned by soaking the electrodes in acetone, cleaned with detergent and polished. These flakes will reduce the electrical conductivity of the electrolyte solution and the electrodes. The surface of the electrode looks relatively smooth without showing any pores.

Metals for electrodes are generally made noble, from gold (Au) and platinum (Pt), which are relatively more resistant to oxidation reactions. In noble metals, the oxide is found in the form of a black crust on the metal surface after oxidation reaction occurs. This crust is easy to clean by sweeping it with a cloth. Oxygen comes from metal oxides which is included via oxidation or functional groups of organic molecules. If the chemical bonding is weaker than the physical bonding, the oxides that have accumulated on the metal surface can be removed from the surface. This situation occurs in precious metals. Resistance to oxidation reactions affects the performance and stability of the electrode.

Preliminary tests were conducted to determine the electrochemical properties of the electrode in reducing Cr3+ and Cr6+. The electric current shows this ability as an electrode response to the applied voltage. The tests were done in electrolyte solutions containing Cr3+ and Cr6+ ions. Cyclic voltammetry measurements are used to provide information about the electrochemical oxidation-reduction reactions that occur in the reaction system and the reversibility of the reactions. 29.

Fig. 2. Voltammogram of solution without electrolyte ion (red line),  electrolyte solution with Cr3+ 1 gL-1 (blue line) and (c) electrolyte solution Cr6+ 1 gL-1.

It can be seen in Figure 2 that the voltammogram of the electrolyte without Cr3+ and Cr6+ did not show peaks (red line). Inversely, the peaks are visible in two other voltammograms, revealing the response of solution containing Cr3+ (and Cr6+ (blue and green lines).  The peak on the voltammograms for measuring electrolytes concentration of Cr3+ and Cr6+ ions is in the anodic current. The scan potential runs from low to high values to generate an anodic current. The opposite direction, cathodic current, shows no peak as expected. It indicates that only reduction reactions occur in electrochemical systems using electrolytes of Cr3+ or Cr6+  and the electrodes. The faradaic process for reducing ions on the surface of the electrode in response to a given potential is specific.

The electroreduction of Cr6+ to Cr3+ occurs at a potential of +1.33 V vs. Standard Hydrogen Electrode (SHE). The potential reduction of Cr3+ to Cr2+ occurs at -0.41 V. In another study the reduction of Cr6+ to Cr3+ at a potential of -0.62 V (vs. Ag/AgCl) 30. The reduction of Cr3+ to Cr2+ was found at -0.74 V (vs Ag/AgCl) 31.

The preliminary study of the Cr6+ reduction current appeared at +0.84 V (vs. Ag) and Cr3+ reduction at +0.74 V (vs. Ag) (Figure 1). The reduction peak shifts from its standard potential are affected by many factors, including the use of electrodes, electrolytes, and the conditions of measurement 32.

The potentials for the electrochemical reactions in which are implicated various Cr(VI)  species were proposed by the other33 as the probable mechanism. The mechanism for concentrations ≥ 25 mM K2CrO4, is given by the following:

HCr2O4-  +   H+   + 3e-      (CrO33-)ads   +   H2O

(CrO33-)ads    +    3HCrO4-  +   H+       2(HCr2O7)ads  +   H2O   +   3e-

2(HCr2O7)ads  +   26H+   +   12e-        4Cr3+    +  14H2O

The H2Cr2O7- species adsorbed on the metal surface by the proposed mechanism is

M(s)   +    HCr2O7-         (-e-)     MI-HCr2O7

The reaction competes with the hydroxyl species of water

M(s)         +    -OH          (-e-)   MI-(OH)

M(s)         +     2H2O     MI(OH2)2+

MI(OH2)2+ +     -OH       MI-(OH)			Where M = Fe, Ni

Fig.  3. (a) Voltammogram of Cr3+ 1 gL-1 in pH varies (b) plot of the relationship of the pH Cr3+ solution and the current peak of the voltammogram

Fig. 4. (a) Voltammogram of Cr6+ 1 gL-1 varies in pH,  (b) plot of the relationship of the pH Cr6+ solution and the current peak of the voltammogram

The Conditions for the measurement

The sensitivity of the measurements is affected by the measurement conditions, i.e., pH, deposition time, and the scan rate. Scan rate and pH gave linear voltammetry measurements of peak currents as reported by Latif et.al.34 Deposition time  affects the ionic deposition on the surface of the electrode35. In general, voltammetry measurements require that the measured faradaic current be slightly larger than the portion of its response to a given potential difference. This condition results in better sensitivity and smaller detection limits.

Optimum pH

Interaction between electrolyte and electrode takes place on the surface 36 pH affects the electrode surface at metal electrodes, which causes the formation of charged ions on the surface, significantly the negative charge of the water molecule. It affects the process at the electrode 37.

The effect of pH was studied in pH 4 – 8 by comparing the voltammogram of the solutions containing Cr3+ and Cr6+. It is expected that more ions are adsorbed at the surface of the electrode in order to produce a maximum and stable peak current for the measurement at suitable pH conditions 38.

Figure 3 shows the change in the response of the Cr3+ ion. The change in current is proportional to the applied potential. It was found that the maximum pH is 6. Figure 4 shows the change in the response of the Cr6+ ion. The change in current is proportional to the applied potential. It was found that the maximum pH is 4. The chromium ion exists in a mixed-valent state 39 with an electric charge depending on the system's pH 40. When the pH of the solution is between 1 and 6, chromium is in equilibrium as it forms HCrO4- and Cr2O72- When the solution is at 6, Cr6+ is in the form of CrO42- ions, while the pH is less than 1 the main species is H2CrO4.  pH 4 was used as the optimum pH for the measurement of Cr6+ because the measurement at pH 4 was more stable than another pH, and the measurement at pH 4 could obtain a higher measurement  concentration 41.

Fig. 5. (a) Voltammogram of Cr3+ 1 gL-1 varies in deposition times. (b)  plot relationship of deposition time and current peak of the voltammogram.

Fig. 6. (a) Voltammogram of Cr6+ 1 gL-1 varies in deposition times, (b) plot relationship of deposition time and  current peak of the voltammogram.

Optimum Scan Rate

The scan rate affects the current peak and width of the voltammogram. The Faraday process, i.e., electrochemical oxidation-reduction reaction, causes a current peak. The change in the width of the voltammogram is due to the diffusion of ions to the electrolyte surface. The effect of scan rate was studied by comparing the voltammogram of the measurement results of Cr3+ and Cr6+ ions at pH conditions and deposition time. The scan rate varied between 100-200 mVs-1. It can be seen from Figures 7 (a) and 8 (a) that the scan rate affects the current peak of Cr3+ and Cr6+ ions. The current was proportional to the electrochemical reaction rate at the electrode surface. The higher the scan rate, the faster the electrolysis reaction speed so that the peak current increases 42. The measurement of Cr3+ at 160 mVs-1 gives the maximum current peak (Figure 7 (a)).  The optimum conditions for measuring Cr3+ were found at pH 6, deposition time of 30 seconds, and a scan rate of 160   mVs-1. The curve is relatively broader compared to the other scan rate. A widening curve interferes with the measurement, especially if there is interference from other ions with adjacent half-wave potentials, which can cause the curves to overlap. There is a decrease in current peak at a higher scan rate (Figure 7(b) because the metal attached to the electrode surface is saturated, so the Cr3+ ions adsorbed on the electrode surface are desorbed.  The highest current for the scan rate of Cr6+ measurement was 140 mVs-1 (Figure 8 (a)). The optimum pH for measuring Cr6+ is 4, deposition time of 60 seconds, and a scan rate of 140 mVs-1. Two maxima exist in the measurement involving the relationship between the scan rate and the current peak.

Fig. 7. (a) Voltammogram of Cr3+ 1 gL-1 in various scan rates (b) Plot of the relationship between scan rate and current peak of the measurement.

Fig. 8. (a) Voltammogram of Cr6+ 1 gL-1 in various scan rates. (b) Plot of the relationship between scan rate and the current peak of the  measurement.

Linear Regression of Cr3+ and Cr6+ Voltammogram

A measurement contains linear and nonlinear parts in a certain measurement range. The linear portion of the measurement is used to quantify the linearity of the measurement. This method is applied to obtain the linearity of Cr3+ and Cr6+ measurement using voltammetry cyclic and the electrode. The linear regression plot was formed from the LSV measurement data for each type of ion. Each type of ion measured varied its concentration in the electrolyte solution, i.e., 0.1; 0.5; 1; 5 and 10 gL-1. The measurements were carried out at pH, deposition, and scan rate to reach maximum current (ip) conditions in anodic current.

Fig. 9. Voltammograms of (a) Cr3+ and (b) Cr6+ Measurements Using Each Optimal Conditions.

The voltammogram for each measurement of Cr3+ and Cr6+ and their concentration variations are shown in Figure 9. The maximum currents or current peaks  (ip) in Figure 9 were plotted for a regression plot.   The regression plot of the Cr3+ and Cr6+ standard solution is presented in Figures 10 and 11.

Figure 10 (a) shows the current peak (ip) obtained at a concentration of 0.1-0.5 gL-1 from Figure 9 (a) as the base to determine the linearity of measurement. Statistical calculations of the plot give 0.8248 for the correlation coefficient (r) and y = 2.3435x + 16.037 for the plot equation. The value of r gives information for the fitness of experimental data and the plot of mathematical equations 43.

Fig. 10. (a) regression plot for ip versus C of Cr3+ ion, (b) linear part of regression plot for Cr3+.

Fig. 11.  (a) regression plot for ip versus C of Cr6+ ion, (b) linear part of regression plot for Cr6+

Some points from the measurement need to be removed to make the fitness of the measurement increase. The last two points, namely C= 0.1 and 0.5, were removed from the regression plot so that a new regression line was obtained with the value of the correlation coefficient (r) = 0.9883 and y = 1.415 x + 23.475 (Figure  9(b)).

It can be seen from Figure 10 (b) that the current increases with increasing Cr3+ concentration. Due to the increasing concentration of Cr3+, the greater the amount of Cr3+ in the solution reduced Cr2+ on the electrode surface.

Figure 11 (a) shows the plot of the current peak versus the concentration of the voltammogram in Figure 9 (b). Regression plot has equation y = 1.5623 x + 8.3061 and  correlation coefficient (r)  is 0.9339. The value is considered to be less linear. The same steps were also taken to produce a more fit plot. Two first points were removed from the measurement. It can be shown from Figure 11 (b) that the plot has equation y = 1.2317 x + 10.945 with the correlation coefficient (r) = 0.9872.

Detection Limit of Measurements.

The limit of detection (LOD) determines the lowest concentration detected. The LODs were calculated from the linear regression plot (Figure 10 (b) and 11(b)) using equations (1) and (2) 44. The measurements were carried out at the optimum pH. The LOD value for Cr3+ and Cr6+ are 2.082 and 2.182 gL-1, respectively. These values indicate that the accurate measurement for Cr3+ and Cr6+ can be obtained from the concentration above those concentrations.

Detection of Cr3+ and Cr6+ in the Mixture

The mixed-valence electrolyte was prepared from the individual electrolyte solution of Cr3+ and Cr6+. The current peaks of the mixture (Figure 11) show different profiles from the others produced by individual measurements of Cr3+ and Cr6+. Voltammograms of the mixed-valence electrolyte of Cr3+ and Cr6+ for 1 and 10 μgL-1 are shown in Figure 12. The Cr6+ has a more negative potential than Cr3+ due to its standard potential 45  and more mobile in an aqueous solution 46.

Fig. 12. Voltammogram of Cr3+, Cr6+ mixed electrolyte with concentration  (a) 1 μgL-1 (b) 10 μgL-1.

Figure 13. The current peaks of Cr3+█ and Cr6+█ in (a) mixture (b) individually electrolyte solution

The ions with a more negative reduction potential will tend to be deposited on the surface of the electrode that has a more positive reduction potential. Cr3+ was reduced first on the electrode surface, followed by a reduction of Cr6+. The more concentration of Cr6+ in the mixture then deposited on the surface of Cr3+ is directly proportional to the current peak of Cr3+. The current of Cr6+ in the mixture decreased due to the presence of the ion, which was deposited on the surface of Cr3+. Cr6+ was reduced to Cr3+ in the reduction potential.

The peak current of the mixture compared to the individual measurement results described in Figure 8, as shown in Figure 13. Plots in Figure 13  reveal that the results of individual measurements at 10 gL-1 are not much different from those of electrolytes in the mixture. The current peak for the Cr6+ ion is not visible at a concentration of 1 gL-1. The peak current is detected as a slight  indentation that slightly changes the slope of the voltammogram. These results indicate that both  spark plug electrodes and voltammetry methods can be used to detect the presence of Cr3+ and Cr6+ in the sample separately, which means it is detected not as total Cr but as its species.

4. Conclusions

X-ray diffraction measurements were carried out to determine the metal elements composed  in the electrode. The diffractogram of the electrode shows the presence of TiNi (2Ɵ  = 26.6,  40.0, 48.8,  54.5)   and overlaps with the diffraction peak of TiCu (2Ɵ = 18.15, 40.0, 54.5). The surface consists of the main structure of the electrode and crack of various sizes. Preliminary tests were conducted to determine the electrochemical properties of the electrode in reducing Cr3+ and Cr6+. The electric current shows this ability as an electrode response to the applied voltage. The tests were done in electrolyte solutions containing Cr3+ and Cr6+ ions. Cyclic voltammetry measurements are used to provide information about the electrochemical oxidation-reduction reactions that occur in the reaction system and the reversibility of the reactions. The reduction of Cr6+ to Cr3+ appeared at +0.84 V (vs Ag/AgCl) and the reduction of Cr3+ to Cr0 at +0.74 V (vs Ag/Ag). The optimum conditions for measuring Cr3+ were found at pH 6, deposition time 30 seconds, and scan rate 160 mVs-1, while for measuring Cr6+ at pH 4, deposition time 60 seconds, and scan rate 140 mVs-1.  The accurate measurement for Cr3+ and Cr6+ can be obtained from LOD value which are The LOD value for Cr3+ and Cr6+ are 2.082 and 2.182 gL-1. There is a slight difference in the individual measurement current and the mixture of Cr3+ - Cr6+  but a good agreement for the oxidation-reduction potential. These results indicate that both  spark plug electrodes and voltammetry methods can be used to detect the presence of Cr3+ and Cr6+ in the sample separately, which means it is detected not as total Cr but as its species. The data analysis shows the feasibility of the electrode Considering the ability of this measurement system to provide a good current response when applying to the ppb range.

5. Acknowledgements

The research/publication of this article was funded by DIPA of Public Service Agency of Universitas Sriwijaya 2021. SP DIPA-042.01.2.400935/2019, On December 05, 2018. In accordance with the Rector’s Decree Number: 0145/UN9/SB3.LP2M.PT/2019.

References

1- 	A. Chait, L. J. den Hartigh, Adipose Tissue Distribution, Inflammation and Its Metabolic Consequences, Including Diabetes and Cardiovascular Disease, Front Cardiovasc Med, 2020, 7, 22.

2- 	E. Sawicka, K. Jurkowska, A. Piwowar, Chromium (III) and chromium (VI) as important players in the induction of genotoxicity – current view, Ann Agric Environ Med, 2021, 28, 1.

3- 	A. M. de O. Lopes, P. R. Chellini, R. A. de Sousa, Cadmium and Chromium Determination in Herbal Tinctures Employing Direct Analysis by Graphite Furnace Atomic Absorption Spectrometry (GF-AAS), Anal Lett, 2020, 53, 2096.

4- 	E. Tsanaktsidou, G. Zachariadis, Titanium and Chromium Determination in Feedstuffs Using ICP-AES Technique, Separations, 2019, 7, 1.

5- 	J. B. Vera, M. C. Bisinoti, C. D. B. Amaral, M. H. Gonzalez, ICP- quadrupole MS for accurate determination of chromium in environmental and food matrices, Environ Nanotechnol Monit Manag, 2021, 15, 100421.

6- 	A. Ait sidi mou, A. Ouarzane, M. El Rhazi, Detection of mercury by a new sensor based CPE modified with extract of takaout plant, Mediterr J Chem, 2016, 5, 514–520.

7- 	N. M. Thanh, N. D. Luyen, T. Thanh Tam Toan, N. Hai Phong, N. Van Hop, Voltammetry Determination of Pb(II), Cd(II), and Zn(II) at Bismuth Film Electrode Combined with 8-Hydroxyquinoline as a Complexing Agent, J Anal Methods Chem, 2019, 2019, 1.

8- 	J. Barón-Jaimez, M. R. Joya, J. Barba-Ortega, Anodic stripping voltammetry – ASV for determination of heavy metals, J Phys Conf Ser, 2013, 466, 012023.

9- 	S. T. Palisoc, R. I. M. Vitto, M. G. Noel, K. T. Palisoc, M. T. Natividad, Highly sensitive determination of heavy metals in water prior to and after remediation using Citrofortunella Microcarpa, Sci Rep, 2021, 11, 1394.

10- 	Y. Liu, J. Liu, J. Liu, W. Gan, B. Ye, Y. Li, Highly sensitive and selective voltammetric determination of dopamine using a gold electrode modified with a molecularly imprinted polymeric film immobilized on flaked hollow nickel nanospheres, Microchim Acta, 2017, 184, 1285.

11- 	J. Lalmalsawmi, Zirlianngura, D. Tiwari, S.-M. Lee, Low cost, highly sensitive and selective electrochemical detection of arsenic (III) using silane grafted based nanocomposite, Environ Eng Res, 2019, 25, 579.

12- 	H. Suyani, I. Rahmi, H. Pardi, Optimization for the Simultaneous Determination of Zinc in Environmental Samples With Calcon by Adsorptive Stripping Voltammetry : Response Surface Methodology, Orient J Chem, 2017, 33, 2060.

13- 	C. M. Welch, R. G. Compton, The use of nanoparticles in electroanalysis: a review, Anal Bioanal Chem, 2006, 384, 601.

14- 	Y. B. Vogel, J. J. Gooding, S. Ciampi, Light-addressable electrochemistry at semiconductor electrodes: redox imaging, mask-free lithography and spatially resolved chemical and biological sensing, Chem Soc Rev, 2019, 48, 3723.

15- 	D. M. Heard, A. J. J. Lennox, Electrode Materials in Modern Organic Electrochemistry, Angew Chem, 2020, 132, 19026.

16- 	G. E. Badea, C. Antal, M. Rosca, A. Setel, Autocatalytic Reduction of Cr(VI) on Platinum Electrode in Acid Solution, Rev Roum Chim, 2012, 57, 29–33.

17- 	W. Liu, G. Liu, S. Xiao, J. Zhang, The Electrochemical Behavior of Cr(II) Ions in NaCl-KCl Melt, Int J Electrochem Sci, 2017, , 1589.

18- 	L. Dunyushkina, A. Pavlovich, A. Khaliullina, Activation of Porous Pt Electrodes Deposited on YSZ Electrolyte by Nitric Acid Treatment, Materials, 2021, 14, 5463.

19- 	M. A. Ehsan, M. H. Suliman, A. Rehman, A. S. Hakeem, A. Al Ghanim, M. Qamar, Fabrication of platinum thin films for ultra-high electrocatalytic hydrogen evolution reaction, Int J Hydrog Energy, 2020, 45, 15076.

20- 	W. Wu, Z. Chen, B. Li, X. Cong, Q. Chen, Mechanical and electrochemical properties of platinum coating by double glow plasma on titanium alloy substrate, Russ J Electrochem, 2013, 49, 76–80.

21- 	J. H. O. J. Wijenberg, A. C. A. de Vooys, R. Kortlever, M. T. M. Koper, Oxidation reactions in chromium(III) formate electrolytes at platinum and at a catalytic mixed metal oxide coating of iridium oxide and tantalum oxide, Electrochimica Acta, 2016, 213, 194–200.

22- 	T. A. F. Appia, L. Ouattara, Electrooxidation of simulated wastewater containing pharmaceutical amoxicillin on thermally prepared IrO2/Ti, Mediterr J Chem, 2021, 11, 172.

23- 	M. Ouakki, A. El Fazazi, M. Cherkaoui, Electrochemical deposition of Zinc on mild steel, Mediterr J Chem, 2019, 8, 30–41.

24- 	A. R. Baqer, A. A. Beddai, M. M. Farhan, B. A. Badday, M. K. Mejbel, Efficient coating of titanium composite electrodes with various metal oxides for electrochemical removal of ammonia, Results Eng, 2021, 9, 100199.

25- 	H. H. Hasan, I. H. A. Badr, H. T. M. Abdel-Fatah, E. M. S. Elfeky, A. M. Abdel-Aziz, Low cost chemical oxygen demand sensor based on electrodeposited nano-copper film, Arab J Chem, 2018, 112, 172.

26- 	A. A. Korda, S. Munawaroh, E. A. Basuki, The Antimicrobial Activity and Characterization of the Cast Titanium Copper Alloys with Variations of Copper Content, IOP Conf Ser Mater Sci Eng, 2019, 547, 012002.

27- 	G. Wilhelm Sievers, K. Anklam, R. Henkel, T. Hickmann, V. Brüser, Corrosion-protection of moulded graphite conductive plastic bipolar plates in PEM electrolysis by plasma processing, Int J Hydrog Energy, 2019, 44, 2435–2445.

28- 	S. Tan, T. Ors, H. Akyildiz, A. F. Kalcioglu, T. Ozturk, Synthesis of FeTi Hydrogen Storage Alloys for Stationary Applications, , In, Proceedings 2nd International Hydrogen Energy Congress and Exhibitio, , Istanbul, Turkey, 2007.

29- 	N. Elgrishi, K. J. Rountree, B. D. McCarthy, E. S. Rountree, T. T. Eisenhart, J. L. Dempsey, A Practical Beginner’s Guide to Cyclic Voltammetry, J Chem Educ, 2018, 95, 197–206.

30- 	Q. Ge, X. Feng, R. Wang, R. Zheng, S. Luo, L. Duan, Y. Ji, J. Lin, H. Chen, Mixed Redox-Couple-Involved Chalcopyrite Phase CuFeS 2 Quantum Dots for Highly Efficient Cr(VI) Removal, Environ Sci Technol, 2020, 54, 8022.

31- 	N. Mehdipour, M. Rezaei, Z. Mahidashti, Influence of glycine additive on corrosion and wear performance of electroplated trivalent chromium coating, Int J Miner Metall Mater, 2020, 27, 544.

32- 	B. Vercelli, S. Crotti, M. Agostini, Voltammetric responses at modified electrodes and aggregation effects of two anticancer molecules: irinotecan and sunitinib, New J Chem, 2020, 44, 18233.

33- 	S. Trasatti, Work function, electronegativity, and electrochemical behaviour of metals, J Electroanal Chem Interfacial Electrochem, 1972, 39, 163–184.

34- 	I. A. Latif, S. H. Merza, Effect of Scan Rate and pH on Determination Amoxilline Using Screen Printed Carbon Electrode Modified with Functionalized Graphene Oxide, Ibn Al-Haitham Jour Pure Appl Sci, 2018, 31, 157–171.

35- 	A. Bobrowski, A. Królicka, R. Bobrowski, Renewable silver amalgam film electrodes in electrochemical stripping analysis—a review, J Solid State Electrochem, 2016, 20, 3217–3228.

36- 	A. T. Kuhn, C. Y. Chan, pH changes at near-electrode surfaces, J Appl Electrochem, 1983, 13, 189.

37- 	A. Goyal, M. T. M. Koper, The Interrelated Effect of Cations and Electrolyte pH on the Hydrogen Evolution Reaction on Gold Electrodes in Alkaline Media, Angew Chem Int Ed, 2021, 60, 13452.

38- 	A. Ganassin, P. Sebastián, V. Climent, W. Schuhmann, A. S. Bandarenka, J. Feliu, On the pH Dependence of the Potential of Maximum Entropy of Ir(111) Electrodes, Sci Rep, 2017, 7, 1246.

39- 	X. Guo, A. Liu, J. Lu, X. Niu, M. Jiang, Y. Ma, X. Liu, M. Li, Adsorption Mechanism of Hexavalent Chromium on Biochar: Kinetic, Thermodynamic, and Characterization Studies, ACS Omega, 2020, 5, 27323.

40- 	A. G. Caporale, A. Violante, Chemical Processes Affecting the Mobility of Heavy Metals and Metalloids in Soil Environments, Curr Pollut Rep, 2016, 2, 15.

41- 	M. Tumolo, V. Ancona, D. De Paola, D. Losacco, C. Campanale, C. Massarelli, V. F. Uricchio, Chromium Pollution in European Water, Sources, Health Risk, and Remediation Strategies: An Overview, Int J Environ Res Public Health, 2020, 17, 5438.

42- 	A. A. Bojang, H. S. Wu, Characterization of Electrode Performance in Enzymatic Biofuel Cells Using Cyclic Voltammetry and Electrochemical Impedance Spectroscopy, Catalysts, 2020, 10, 782.

43- 	D. Chicco, M. J. Warrens, G. Jurman, The coefficient of determination R-squared is more informative than SMAPE, MAE, MAPE, MSE and RMSE in regression analysis evaluation, PeerJ Comput Sci, 2021, 7, e623.

44- 	R. Ismail, H. Y. Lee, N. A. Mahyudin, F. Abu Bakar, Linearity study on detection and quantification limits for the determination of avermectins using linear regression, J Food Drug Anal, 2014, 22, 407.

45- 	M.-J. Jorge, M.-C. Nilson, H.-R. Aracely, F. Machuca-Martínez, Data on the removal of metals (Cr, Cr, Cd, Cu, Ni, Zn) from aqueous solution by adsorption using magnetite particles from electrochemical synthesis, Data Brief, 2019, 24, 103956.

46- 	R. Heydarzadeh, A. A. Ghadimkhani, A. Torabian, An improvement of Cr+6 removal by the reduction to Cr+3 in Birjand ground water treatment, , In, Water Pollution VIII: Modelling, Monitoring and Management, , WIT Press: Bologna, Italy, 2006, p, 195.